Clave 16 PM Checklist:

- Inspect display, touchpad, USB port, printer port, safety valve, and power switch. Note any issues and order parts for another visit. Turn off unit and unplug.

- Drain distilled water tank and waste water tank. Clean distilled water tank. Replace distilled water tank filter.

- Check door handle and door operation. Adjust if necessary.

- Clean chamber and inside door. You can use ‘Comet’ type cleaner (not supplied) in the chamber. Use as little as possible and rinse out several times with water. Use a Q-tip to clean out the holes in the back wall. Replace door gasket. Replace mesh filter inside chamber.

- Check both drain ports for leaks. Replace outside section only or both if leaking. Lube ‘O’ rings.

- Remove side and top covers. Inspect wire/cable connections to board and assemblies/components. Insure that none are loose or discoloured from overheating. Correct or replace as necessary.

- Inspect tubing for hardness or discoloration. Replace any that are suspect. It should not be necessary to replace all tubing.

- Replace check valve at the waste water tank. Keep the correct orientation as original.

- Replace battery (CR2032) every 2 years, obtain locally.

- Run a short cycle (Unwrapped 134C) with a Class 5 test strip. Check for leaks. After a successful test with no leaks, re-install sides and top cover.

- Reset the maintenance counter and date.

- Hold the ‘P’ button while powering on. Release it after entering the S1 screen. Press the start button to enter that screen. Press the ‘P’ button to select the number to change. Press the temp button to change the number. Set the cycle count to 500 or the number of cycles the customer will do in a year and the date to 1 yr in the future. E88 will come up when either value is reached. Press the start button to save.

- Return unit to user.

PM kit parts:   Retail prices for parts, tube prices per piece are below.

60cm silicon tube, 4 x 8mm	p/n 90012027	($12.42/m)

175cm silicon tube, 6 x 10mm	p/n 90012028	($18.63/m)

175cm silicon tube, 6 x 12mm	p/n 90012029	($24.84/m)

35cm silicon tube, 16 x 20mm	p/n 90012031	($27.50/m)

Door gasket			p/n 90012034	($45.00)

Check valve			p/n 90012063	($16.50)

Drain port (x2)			p/n 90012047	($16.00ea)

Chamber mesh filter		p/n 90012050	($7.00)

Distilled water tank filter	p/n 90012049	($7.00)

Tube ring 8.3mm		p/n 90012026	($5.50)

Tube ring 12.7mm		p/n 90012032	($5.50)

Class 5 test strip		(not supplied)

Nylon Tywraps			(20 supplied)

Cup washer 			p/n 90012113	($1.00)

Panel screw 			p/n 90012114	($1.00)

Clave 16 Tube Lengths:

Estimates only. Remove old tube and line up against new tube (add a little if original tube was tight) and cut square. You can always trim a longer tube. You can’t add to a shorter tube. Start with the longest tube first. There is a few extra cm. of tube in the PM kit. Retail prices are for the piece.

4x8mm: (60cm)

From Tee to pressure sensor 	11 cm	($1.37)

From Tee to pressure lock 	4.5 cm	($0.56)

From Tee to back of chamber 	40 cm	($4.96)

6x10mm: (175cm)

From distilled tank to pump 	35 cm	($6.52)

From distilled tank to rear drain 40 cm	($7.45)

From water valve to chamber 	40 cm	($7.45)

From check valve to waste tank 10 cm	($1.86)

From condenser to waste tank 	40 cm	($7.45)

From waste tank to front drain 	8.5 cm	($1.58)

6x12mm: (170cm)

From chamber to safety valve 	50 cm	($12.42)

From pump to steam generator 	18 cm	($4.47)

From steam gen. to chamber 	23 cm	($5.71)

From chamber to steam valve 	19 cm	($4.72)

From steam valve to cross 	3.5 cm	($0.87)

From cross to condenser 	6.5 cm	($1.62)

From cross to water valve 	18 cm	($4.47)

From cross to check valve 	28 cm	($6.96) check valve pointed to cross

16x20mm: (35cm)

From waste tank to rear vent 	35 cm	($9.63)

Use the parts from the kit that are needed and re-stock the kit by ordering just the parts used. In most cases, the tubing will not be fully used.