Clave 23 PM Checklist:

- Inspect display, touchpad, USB port, printer port, safety valve, and power switch. Note any issues and order parts for another visit. Turn off unit and unplug.

- Drain distilled water tank and waste water tank. Clean distilled water tank. Replace distilled water tank filter.

- Check door handle and door operation. Adjust if necessary.

- Clean chamber and inside door. You can use ‘Comet’ type cleaner (not supplied) in the chamber. Use as little as possible and rinse out several times with water. Use a Q-tip to clean out the holes in the back wall. Replace door gasket. Replace mesh filter inside chamber.

- Check both drain ports for leaks. Replace outside section only or both if leaking. Lube ‘O’ rings.

- Remove side and top covers. Inspect wire/cable connections to board and assemblies/components. Insure that none are loose or discoloured from overheating. Correct or replace as necessary.

- Inspect tubing for hardness or discoloration. Replace any that are suspect. It should not be necessary to replace all tubing.

- Replace check valve at the waste water tank. Keep the correct orientation as original.

- Replace battery (CR2032) every 2 years, obtain locally.

- Run a short cycle (Unwrapped 134C) with a Class 5 test strip. Check for leaks. After a successful test with no leaks, re-install sides and top cover.

- Reset the maintenance counter and date.

- Hold the up and down buttons while powering on. Let go when you enter the interface. Scroll to ‘Maintenance’ and highlight it. Press Select button. Use the select button to change the ‘End’ date to 1 yr in the future. Change the cycle count to 500 or the number of cycles the customer will do in a year. Use the up/down buttons to change the values. When finished, press the select button and power off. E88 or ‘wrench’ icon will come up again when either value is reached.

- Return unit to user.

PM kit parts:  Retail prices for the parts, tube prices per piece are below.

70cm silicon tube, 4 x 8mm	p/n 90012027	($12.42/m)

230cm silicon tube, 6 x 10mm	p/n 90012028	($18.63/m)

170cm silicon tube, 6 x 12mm	p/n 90012029	($24.84/m)

40cm silicon tube, 16 x 20mm	p/n 90012031	($27.50/m)

Door gasket			p/n 90012045	($45.00)

Check valve			p/n 90012063	($16.50)

Drain port (x2)			p/n 90012047	($16.00ea)

Chamber mesh filter		p/n 90012050	($7.00)

Distilled water tank filter	p/n 90012049	($7.00)

Tube ring 8.3mm		p/n 90012026	($5.50)

Tube ring 12.7mm		p/n 90012032	($5.50)

Class 5 test strip		(not supplied)

Nylon Tywraps			(20  supplied)

Cup washer 			p/n 90012113	($1.00)

Panel screw 			p/n 90012114	($1.00)

Clave 23 Tube Lengths:

Estimates only. Remove old tube and line up against new tube (add a little if original tube was tight) and cut square. You can always trim a longer tube. You can’t add to a shorter tube. Start with the longest tube first. There is a few extra cm. of tube in the PM kit. Retail prices for the piece.

4x8mm:

From Tee to pressure sensor 	16 cm	($1.99)

From Tee to pressure lock 	5.5 cm	($0.68)

From Tee to back of chamber 	46 cm	($5.71)

6x10mm:

From distilled tank to pump 	35 cm	($6.52)

From distilled tank to front drain 60 cm	($11.18)

From Tee to check valve 	30 cm	($5.59) check valve pointed to Tee

From check valve to waste tank 10 cm	($1.86)

From condenser to waste tank 	43 cm	($8.01)

From waste tank to front drain 	50 cm	($9.32)

6x12mm:

From chamber to safety valve 	38 cm	($9.44)

From pump to steam generator 	16.5 cm	($4.10)

From steam gen. to chamber 	23 cm	($5.71)

From chamber to steam valve 	20 cm	($4.97)

From steam valve to cross 	3 cm	($0.75)

From Tee to condenser 		6 cm	($1.50)

From Tee to water valve 	18 cm	($4.47)

From water valve to chamber 	40 cm	($9.94)

16x20mm:

From waste tank to rear vent 	40 cm	($11.00)

Use the parts from the kit that are needed and re-stock the kit by ordering just the parts used. In most cases, the tubing will not be fully used.